Summary of Studies

Performed on

###### Sklar ®

Disinfectant & Soak

US EPA Registration No. 70144-1-31118

Revised June 2005

Description of Products

### Sklar Disinfectant / Soak solution are trademarked brand names belonging to Sklar Instrument Company. Sklar Disinfectant and Soak are produced from a combination of a number of solvents, sequestering agents, detergents and other ingredients acting in synergy with three distinct antimicrobial active agents.  These products may be applied directly to surfaces made of plastic, Plexiglas, baked on painted surfaces, acrylics, stainless steel and glass.

Sklar Disinfectant / Soak solution are a highly effective decontaminant / disinfectant intended for the decontamination and disinfection of medical / surgical / dental / ophthalmic equipment surfaces, instruments and devices.  They are particularly useful in acute patient care areas, surgical suites, GI labs, kidney dialysis facilities and other areas where the control of cross-contamination is important.

Studies Performed & Results

Having demonstrated biocidal effectiveness when tested with a three (3) minute exposure time against the following representative families of microorganisms, the Sklar Disinfectant / Soak solution formula is proven to be bactericidal, tuberculocidal, virucidal and fungicidal within three (3) minutes.  The following study summaries include:

Microbiocidal Efficacy

Mycobacterium tuberculosis (TB)

Staphylococcus aureus

Pseudomonas aeruginosa

Salmonella choleraesuis

Hepatitis b virus (HBV)

Hepatitis c virus (HCV)

Herpes Simplex virus (HSV)

Human Immunodeficiency virus (HIV-1)

Poliovirus

Rhino Virus

Influenza Virus

Clostridium difficile

Vancomycin Resistant Enterococcus (VRE)

Methicillin Resistant Staphylococcus Aureus  (MRSA)

Toxicity / Irritation Studies

Oral Toxicity

Inhalation Toxicity

Dermal Toxicity / Irritation / Sensitization

Ocular Irritation

Mycobacterium tuberculosis bovis BCG (tb) studies:

“A Quantitative Suspension Test for Determining Tuberculocidal Activity”.

The results show Sklar Disinfectant / Soak solution to be an effective tuberculocidal agent when used as directed.

## “AOAC Tuberculocidal Study”

Three lots of Sklar Disinfectant / Soak solution exhibited no growth of mycobacterium bovis BCG when used as directed.

“AOAC Confirmative Tuberculocidal Activity of Sklar Disinfectant / Soak solution”

Sklar Disinfectant / Soak solution killed 100% of the Mycobacterium bovis BCG label cylinders within three (3) minutes at 20C

“Sklar Disinfectant / Soak  solution  vs. Mycobacterium bovis BCG in a Rate of Kill Suspension Study”

In a suspension over a period of three separate tests, Sklar Disinfectant / Soak solution consistently killed 100% Mycobacterium bovis BCG

Bactericidal Studies:

“Staphylococcus aureus AOAC Germicidal Spray Test”

Diluted Sklar Disinfectant / Soak solution (worst case solution with minimum production concentrations of active ingredients and Isopropanol) passed the AOAC Spray Test within 3 minutes when tested against S. aureus at 20C.

“AOAC Use-Dilution Test: (Confirmatory) Evaluation of the Efficacy of the Sklar Disinfectant / Soak solution formula against S. aureus”

Two lots of Sklar Disinfectant / Soak solution, undiluted demonstrated no growth on any of the carriers in the primary subculture when tested against S. aureus.  Under the conditions of this study, Sklar Disinfectant / Soak solution is considered cidal with three (3) minute contact time at 2 º C. against S. aureus.

“Pseudomonas aeruginosa AOAC Germicidal Spray Test”

Diluted Sklar Disinfectant / Soak solution (worst case solution with minimal production concentrations of active ingredients and Isopropanol) passed the AOAC Spray Test when tested against P. aeruginosa at 20C. with three (3) minute contact time

“AOAC Use-Dilution Test:” (Confirmatory) Evaluation of the Efficacy of the Sklar Disinfectant / Soak solution formula against P. aeruginosa.”

Two lots of Sklar Disinfectant / Soak solution, undiluted demonstrated no growth on any of the carries in the primary subculture when tested against P. aeruginosa. Under the conditions of this study, Sklar Disinfectant / Soak solution is considered cidal against P. aeruginosa within three (3) minutes at 20º C.

“Salmonella choleraesuis AOAC Germicidal Spray Test “

Diluted Sklar Disinfectant / Soak solution (worst case solution with minimum production concentrations of active ingredients passed the AOAC Spray Test when tested against S. choleraesuis at 20C. with a three (3) minute contact time

“AOAC Use-Dilution Test:” (Confirmatory) Evaluation of the Efficacy of the Sklar Disinfectant / Soak solution formula against S. choleraesuis.”

Two lots of Sklar Disinfectant / Soak solution, undiluted demonstrated no growth on any of the carries in the primary subculture when tested against. S. choleraesuis . Under the conditions of this study, Sklar Disinfectant / Soak solutionis considered cidal against S. choleraesuis at 20 C .with a three (3) minute contact time

“AOAC Use-Dilution Test” Sklar Disinfectant / Soak solution vs. S. aureus, P.aeruginosa and S. choleraesuis

Three lots of Sklar Disinfectant / Soak solution showed no growth when tested for three (3) minutes at 20C.

“Bactericidal Activity of Sklar Disinfectant / Soak solution in a Stainless Steel Cylinder Test and in Suspension”

These studies were designed to determine how quickly bacteria were killed on stainless steel surfaces, and while in suspension at 20C.  The studies were designed to determine whether Sklar Disinfectant / Soak solution actually killed the bacteria, rather than simply inhibit its replication.  The studies determined that Sklar Disinfectant / Soak solution killed the bacteria.

“Quantitative Analysis of the Kill or Removal of Bacteria in Heavy Organic Soil from Surfaces by Sklar Disinfectant / Soak solution” (custom-designed study)

This study was designed to measure the extent to which Sklar Disinfectant / Soak solution could kill S. aureaus, P. aeruginosa and S. choleraesuis dried in concentrations up to 20% calf serum on glass slide surfaces.  At the conclusion no surviving bacteria were found in the Sklar Disinfectant / Soak solution drained from the glass cover slips

“Clostridium difficile AOAC Germicidal Spray Test in the presence of heat-inactivated horse serum to achieve a 5% organic soil load.”

Two lots of Sklar Disinfectant / Soak solution passed the AOAC Spray Test within three (3) minutes when tested against C. difficile at 20C.

“Methicillin Staphylococcus Aureus (MRSA) Germicidal Spray Test”

Two lots of Sklar Disinfectant / Soak solution passed the Germicidal Spray test with a three (3) minutes contact time at 20C.

“Vancomycin Resistant Enterococcus  (VRE) Germicidal Spray Test”

Two lots of Sklar Disinfectant / Soak solution passed the Germicidal Spray test with a three (3) minute contact time at 20C.

##### Fungicidal Activity Studies

“Fungicidal Activity in a Stainless Steel Cylinder Use-Dilution Test”

Sklar Disinfectant / Soak solution killed Trichophyton mentagrophytes in suspension within 30 seconds at 20C. and within 1 minute on stainless steel surfaces at 20C

“AOAC Fungicidal Test”

Three lots of Sklar Disinfectant / Soak solution exhibited no growth of Trichophyton mentagrophytes at three (3) minutes at 20C

#### Virucidal Activity Studies

“Effectiveness of  Sklar Disinfectant / Soak solution to Inactivate Herpes Simplex Virus Type 1”

Sklar Disinfectant / Soak solution inactivated HSV with a 30 seconds contact time at 20C

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Polio Virus Type 1”

Sklar Disinfectant / Soak solution inactivated Polio type 1 with a three (3) minute contact time at 20C

“Effectiveness of Sklar Disinfectant / Soak solution Liquid (Suspension) to Inactivate Polio Virus Type 1”

Sklar Disinfectant / Soak solution inactivated Polio Virus with three (3) minutes contact time at 20C

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Hepatitis C Virus using Bovine Diarrhea Virus”.   Sklar Disinfectant / Soak solution inactivated HCV with a three (3) minutes contact time at 20º C.

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Hepatitis B Virus using Duck Hepatitis B Virus”.  Sklar Disinfectant / Soak solution inactivated HBV with three (3) minutes contact time at 20º C.

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Rhinovirus”

Sklar Disinfectant / Soak solution inactivated Rhinovirus with 30 seconds contact time at 20º C.

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Influenza virus”

Sklar Disinfectant / Soak solution inactivated the Influenza virus with 30 seconds contact time at 20º C.

“Effectiveness of Sklar Disinfectant / Soak solution Spray to Inactivate Human Immunodeficiency Virus Type 1 (HIV-1, associated with AIDS)”

Sklar Disinfectant / Soak solution inactivated the HIV-1 virus with a 30 seconds contact time at 20º C.

Toxicity / Irritation Studies

The overall results of the listed studies show Sklar Disinfectant / Soak solution Disinfectant solution to be safe for use when used as directed.

“Oral Toxicity”

Sklar Disinfectant / Soak solution was tested for potential acute oral toxicity in accordance with the procedure outlined in the Pesticide Assessment Guideline, US EPA.  No signs of toxicity were exhibited during the 14-day observation period of this study.

Based on the results, the acute oral toxicity LD 50 of Sklar Disinfectant / Soak solution is greater than 5g/kg of body weight.  This product is not considered an oral toxin.

“Acute Inhalation Toxicity”

An acute Inhalation Toxicity Study was conducted to determine the potential for Sklar Disinfectant / Soak solution to produce toxicity via the inhalation route at an exposure level of 2.0 mg/L. Based on the results, the single exposure Acute Inhalation LD 50 of the test solution is greater than 2.13 mg/L. The test results show this product to be categorized in the safest category (Cat IV)

for chemical pesticides and is not a primary inhalation toxin.

” Primary Dermal Irritation”

Sklar Disinfectant / Soak solution was tested for potential dermal irritation in accordance with the procedures outlined in the Pesticide Assessment Guidelines, US EPA.  Sklar Disinfectant / Soak solution exhibited slight reversible redness during the observation period.  Based on the results, Sklar Disinfectant / Soak solution may produce some reversible, slight skin irritation if directly applied to skin.

“Acute Dermal Toxicity Study”

Sklar Disinfectant / Soak solution was tested to evaluate its potential dermal toxicity.  The specimens did not exhibit any signs of toxicity during the 14-day observation period following exposure. Based on the results of this study the LD 50 is greater that 2.0 g/kg of body weight and is non-toxic to skin.

“Primary Eye Irritation Study”

New Zealand Albino Rabbits weighing 2.0-3.0 kg were employed to evaluate the potential irritant effects of Sklar Disinfectant / Soak solution on the eye mucosa.  Based on the criteria outlined in Grades for Ocular Lesions: Pesticide Assessment Guidelines, US EPA,

Sklar Disinfectant / Soak solution produced some slight reversible eye irritation.  The results indicate that Sklar Disinfectant / Soak solution may produce reversible moderate eye irritation when instilled directly into the eye.

Additional technical information may be obtained from:

Micro-Scientific Industries, Inc.

Technical –Scientific Support Group

1-888-253-2536.

Sklar Disinfectant / Soak solutions are produced following strict FDA Good Manufacturing Practices standards and are registered by the US EPA.